## Druck
uuid | Ident-Nummer | Messbereich von | Messbereich bis | Messbereich Einheit | absolut/ relativ | Ausgabebereich von | Ausgabebereich bis | Ausgabebereich Einheit | Kennlinie Steigung _ Sensitivity | Kennlinie Offset _ Bias | Sensitivity Uncertainty | Sensitivity Uncertainty Unit | Sensitivity Uncertainty Comment | Sensitivity Uncertainty Keywords | Bias Uncertainty | Bias Uncertainty Unit | Bias Uncertainty Comment | Bias Uncertainty Keywords | Linearity Uncertainty | Linearity Uncertainty Unit | Linearity Uncertainty Comment | Linearity Uncertainty Keywords | Hysteresis Uncertainty | Hysteresis Uncertainty Unit | Hysteresis Uncertainty Comment | Hysteresis Uncertainty Keywords | Messprinzip | Hersteller | Hersteller P_ID | Bezeichnung | Seriennummer | Grenzfrequenz (-3 dB) | Verantwortlicher WiMi | Aufbewahrungsort | Datenblatt Link | Kalibrationsprotokoll Link | letzte Prüfung/ Kalibration | Zubehör | Bemerkung | Datasheet FST-Link
0184ebd9-988b-7bb9-aa19-1b8573bd0a50 | D001 | 0 | 0.2 | bar |  | 0 | 10 | V | 0.02 | 0 | [16 empty] | Piezoresistiv | Keller |  | PD-23/8666.1-0.2 | 83714 | 1 kHz | Rexer | Lager Messtechnik | [5 empty] | S:\Rexer\02_Datenblaetter\datenblatt_080300_Keller_Drucksensor_Serie23u25_200bar.pdf
0184ebd9-988b-7bb9-aa48-34144ed09467 | D002 | 0 | 0.2 | bar |  | 0 | 10 | V | 0.02 | 0 | [16 empty] | Piezoresistiv | Keller |  | PD-23/8666-0.2 | 80876 | 1 kHz | Rexer | Hydropulser Schrank | [5 empty] | S:\Rexer\02_Datenblaetter\datenblatt_080300_Keller_Drucksensor_Serie23u25_200bar.pdf
0184ebd9-988b-7bb9-aab3-4241ffa5be84 | D003 | 0 | 5 | bar |  | 0 | 10 | V | 0.5 | 0 | 0.005 | %MV/°C | <= typical: 0.005 %MV/°C (measurement value), maximal: 0.02 %MV/°C (reference temperature according to datasheet most probably at 25°C) | relative_temperature_ambient;relative_mv | 0.00025 | bar/°C | <= typical: 0.005 %FS/°C, maximal: 0.01 %FS/°C (reference temperature according to datasheet most probably at 25°C) | relative_temperature_ambient;absolute | 0.01 | bar | typical: <= +- 0.2 %FS, maximal: <= +- 0.5 %FS | absolute | 0.01 | bar | typical: <= +- 0.2 %FS, maximal: <= +- 0.5 %FS | absolute | Piezoresistiv | Keller |  | PA-23/8465-5 | 74274 | 1 kHz | Rexer | Hydropulser Schrank | [5 empty] | S:\Rexer\02_Datenblaetter\datenblatt_080300_Keller_Drucksensor_Serie23u25_200bar.pdf
0184ebd9-988b-7bb9-aace-33be8c477d2a | D004 | 0 | 5 | bar |  | 0 | 10 | V | 0.5 | 0 | 0.005 | %MV/°C | <= typical: 0.005 %MV/°C (measurement value), maximal: 0.02 %MV/°C (reference temperature according to datasheet most probably at 25°C) | relative_temperature_ambient;relative_mv | 0.00025 | bar/°C | <= typical: 0.005 %FS/°C, maximal: 0.01 %FS/°C (reference temperature according to datasheet most probably at 25°C) | relative_temperature_ambient;absolute | 0.01 | bar | typical: <= +- 0.2 %FS, maximal: <= +- 0.5 %FS | absolute | 0.01 | bar | typical: <= +- 0.2 %FS, maximal: <= +- 0.5 %FS | absolute | Piezoresistiv | Keller |  | PAA-23/5bar/8465 | 81248 | 1 kHz | Rexer | Hydropulser Schrank | [5 empty] | S:\Rexer\02_Datenblaetter\datenblatt_080300_Keller_Drucksensor_Serie23u25_200bar.pdf
0184ebd9-988b-7bb9-ab3a-32835e001a30 | D05 | 0 | 6.9 | bar |  | 0 | 10 | V | 0.69 | 0 | [16 empty] | Piezoresistiv | setra |  | 290E | 135530 |  | Rexer | Hydropulser Schrank
0184ebd9-988b-7bb9-abbd-c811ee258dd1 | D06 | 0 | 10 | bar |  | 0 | 10 | V | 1 | 0 | 0.005 | %MV/°C | <= typical: 0.005 %MV/°C (measurement value), maximal: 0.02 %MV/°C (reference temperature according to datasheet most probably at 25°C) | relative_temperature_ambient;relative_mv | 0.0005 | bar/°C | <= typical: 0.005 %FS/°C, maximal: 0.01 %FS/°C (reference temperature according to datasheet most probably at 25°C) | relative_temperature_ambient;absolute | 0.02 | bar | typical: <= +- 0.2 %FS, maximal: <= +- 0.5 %FS | absolute | 0.02 | bar | typical: <= +- 0.2 %FS, maximal: <= +- 0.5 %FS | absolute | Piezoresistiv | Keller |  | PAA-23/8465.1-10 | 22563 | 1 kHz | Rexer | Lager Messtechnik | [5 empty] | S:\Rexer\02_Datenblaetter\datenblatt_080300_Keller_Drucksensor_Serie23u25_200bar.pdf
0184ebd9-988b-7bb9-abdd-10d380488b52 | D07 | 0 | 15 | bar |  | 0 | 10 | V | 1.5 | 0 | 0.005 | %MV/°C | <= typical: 0.005 %MV/°C (measurement value), maximal: 0.02 %MV/°C (reference temperature according to datasheet most probably at 25°C) | relative_temperature_ambient;relative_mv | 0.00075 | bar/°C | <= typical: 0.005 %FS/°C, maximal: 0.01 %FS/°C (reference temperature according to datasheet most probably at 25°C) | relative_temperature_ambient;absolute | 0.03 | bar | typical: <= +- 0.2 %FS, maximal: <= +- 0.5 %FS | absolute | 0.03 | bar | typical: <= +- 0.2 %FS, maximal: <= +- 0.5 %FS | absolute | Piezoresistiv | Keller |  | PAA-23/8465-15 | 5155 | 1 kHz | Rexer | Hydropulser Schrank | [5 empty] | S:\Rexer\02_Datenblaetter\datenblatt_080300_Keller_Drucksensor_Serie23u25_200bar.pdf
0184ebd9-988b-7bb9-ac29-7d6e80d0ceeb | D08 | 0 | 40 | bar |  | 0 | 10 | V | 4 | 0 | [16 empty] | Piezoresistiv | vibro-meter |  | PL 205, 423-205-000-041 | 135 |  | Rexer | Hydropulser Schrank
0184ebd9-988b-7bb9-ac58-3d794672b694 | D09 | 0 | 40 | bar |  | 0 | 10 | V | 4 | 0 | [16 empty] | Piezoresistiv | vibro-meter |  | PL 205, 423-205-000-041 | 138 |  | Rexer | Hydropulser Schrank
0184ebd9-988b-7bb9-acb9-50cd802c4b3b | D10 | 0 | 100 | bar |  | 0 | 10 | V | 10 | 0 | 0.005 | %MV/°C | <= typical: 0.005 %MV/°C (measurement value), maximal: 0.02 %MV/°C (reference temperature according to datasheet most probably at 25°C) | relative_temperature_ambient;relative_mv | 0.005 | bar/°C | <= typical: 0.005 %FS/°C, maximal: 0.01 %FS/°C (reference temperature according to datasheet most probably at 25°C) | relative_temperature_ambient;absolute | 0.2 | bar | typical: <= +- 0.2 %FS, maximal: <= +- 0.5 %FS | absolute | 0.2 | bar | typical: <= +- 0.2 %FS, maximal: <= +- 0.5 %FS | absolute | Piezoresistiv | Keller |  | PA-23/8465-100 | 19175 | 1 kHz | Rexer | Hydropulser Schrank | [5 empty] | S:\Rexer\02_Datenblaetter\datenblatt_080300_Keller_Drucksensor_Serie23u25_200bar.pdf
0184ebd9-988b-7bb9-acf1-3eeb9f37f842 | D11 | 0 | 100 | bar |  | 0 | 10 | V | 10 | 0 | 0.005 | %MV/°C | <= typical: 0.005 %MV/°C (measurement value), maximal: 0.02 %MV/°C (reference temperature according to datasheet most probably at 25°C) | relative_temperature_ambient;relative_mv | 0.005 | bar/°C | <= typical: 0.005 %FS/°C, maximal: 0.01 %FS/°C (reference temperature according to datasheet most probably at 25°C) | relative_temperature_ambient;absolute | 0.2 | bar | typical: <= +- 0.2 %FS, maximal: <= +- 0.5 %FS | absolute | 0.2 | bar | typical: <= +- 0.2 %FS, maximal: <= +- 0.5 %FS | absolute | Piezoresistiv | Keller |  | PA-23/8465-100 | 74905 | 1 kHz | Rexer | Hydropulser Schrank | [5 empty] | S:\Rexer\02_Datenblaetter\datenblatt_080300_Keller_Drucksensor_Serie23u25_200bar.pdf
0184ebd9-988b-7bb9-ad26-ecdae251dc50 | D12 | 0 | 100 | bar |  | 0 | 10 | V | 10 | 0 | 0.005 | %MV/°C | <= typical: 0.005 %MV/°C (measurement value), maximal: 0.02 %MV/°C (reference temperature according to datasheet most probably at 25°C) | relative_temperature_ambient;relative_mv | 0.005 | bar/°C | <= typical: 0.005 %FS/°C, maximal: 0.01 %FS/°C (reference temperature according to datasheet most probably at 25°C) | relative_temperature_ambient;absolute | 0.2 | bar | typical: <= +- 0.2 %FS, maximal: <= +- 0.5 %FS | absolute | 0.2 | bar | typical: <= +- 0.2 %FS, maximal: <= +- 0.5 %FS | absolute | Piezoresistiv | Keller |  | PA-23/8465-100 | 88622 | 1 kHz | Rexer | Hydropulser Schrank | [5 empty] | S:\Rexer\02_Datenblaetter\datenblatt_080300_Keller_Drucksensor_Serie23u25_200bar.pdf
0184ebd9-988b-7bb9-ad69-7b110d486b63 | D13 | 0 | 100 | bar |  | 0 | 10 | V | 10 | 0 | 0.005 | %MV/°C | <= typical: 0.005 %MV/°C (measurement value), maximal: 0.02 %MV/°C (reference temperature according to datasheet most probably at 25°C) | relative_temperature_ambient;relative_mv | 0.005 | bar/°C | <= typical: 0.005 %FS/°C, maximal: 0.01 %FS/°C (reference temperature according to datasheet most probably at 25°C) | relative_temperature_ambient;absolute | 0.2 | bar | typical: <= +- 0.2 %FS, maximal: <= +- 0.5 %FS | absolute | 0.2 | bar | typical: <= +- 0.2 %FS, maximal: <= +- 0.5 %FS | absolute | Piezoresistiv | Keller |  | PA-23/8465-100 | 88623 | 1 kHz | Rexer | Hydropulser Schrank | [5 empty] | S:\Rexer\02_Datenblaetter\datenblatt_080300_Keller_Drucksensor_Serie23u25_200bar.pdf
0184ebd9-988b-7bb9-ad96-201f9b722b5a | D14 | 0 | 100 | bar |  | 0 | 10 | V | 10 | 0 | [16 empty] | Piezoresistiv | vibro-meter |  | PL 207, 423-207-000-041 | 229 |  | Rexer | Hydropulser Schrank
0184ebd9-988b-7bb9-adfa-9d9be1ca8ca9 | D15 | 0 | 100 | bar |  | 0 | 10 | V | 10 | 0 | [16 empty] | Piezoresistiv | vibro-meter |  | PL 207, 423-207-000-041 | 231 |  | Rexer | Hydropulser Schrank
0184ebd9-988b-7bb9-ae37-8468fafa941a | D16 | 0 | 100 | bar |  | 0 | 10 | V | 10 | 0 | [16 empty] | Piezoresistiv | vibro-meter |  | PL 207, 423-207-000-041 | 233 |  | Rexer | Hydropulser Schrank
0184ebd9-988b-7bb9-ae4b-17bd29feb36c | D17 | 0 | 100 | bar |  | 0 | 10 | V | 10 | 0 | [16 empty] | Piezoresistiv | vibro-meter |  | PL 207, 423-207-000-041 | 234 |  | Rexer | Hydropulser Schrank
0184ebd9-988b-7bb9-aea5-aa762bbe9261 | D18 | 0 | 100 | bar |  | 0 | 10 | V | 10 | 0 | [16 empty] | Piezoresistiv | vibro-meter |  | PL 207, 423-207-000-041 | 272 |  | Rexer | Hydropulser Schrank
0184ebd9-988b-7bb9-aeef-ad43c3c02f9d | D19 | 0 | 160 | bar |  | 0 | 10 | V | 16 | 0 | [16 empty] | Piezoresistiv | TransInstruments |  | BHL-4201-00-03 | L423874 |  | Rexer | Hydropulser Schrank
0184ebd9-988b-7bb9-af08-d1e28e2dee91 | D20 | 0 | 160 | bar |  | 0 | 10 | V | 16 | 0 | [16 empty] | Piezoresistiv | TransInstruments |  | BHL-4201-00-03 | L423875 |  | Rexer | Hydropulser Schrank
0184ebd9-988b-7bb9-af6c-03ec06d354cc | D21 | 0 | 200 | bar |  | 0 | 10 | V | 20 | 0 | [16 empty] | Piezoresistiv | Keller |  | PA-25TAB/80087 | 60034 | 1 kHz | Rexer | Lager Messtechnik
0184ebd9-988b-7bb9-afa9-4ad47c116bca | D22 | 0 | 200 | bar |  | 0 | 10 | V | 20 | 0 | [16 empty] | Piezoresistiv | Keller |  | PA-25TAB/80087 | 60035 | 1 kHz | Rexer | Lager Messtechnik
0184ebd9-988b-7bb9-afc2-f3ecf01ec673 | D23 | 0 | 200 | bar |  | 0 | 10 | V | 20 | 0 | [16 empty] | Piezoresistiv | Keller |  | PA-25TAB/80087 | 60036 | 1 kHz | Rexer | Lager Messtechnik
0184ebd9-988b-7bb9-b035-53c4df4ceaa6 | D24 | 0 | 250 | bar |  | 0 | 10 | V | 25 | 0 | [16 empty] | Piezoresistiv | vibro-meter |  | PL 209, 423-209-000-041 | 250 |  | Rexer | Hydropulser Schrank
0184ebd9-988b-7bb9-b04a-915584cbd834 | D25 | 0 | 300 | bar |  | 0 | 10 | V | 30 | 0 | [16 empty] | DMS | Burster |  | 8210-300, 210096 | 43484 | 1 kHz | Rexer | Hydropulser Schrank
0184ebd9-988b-7bb9-b0f3-1542bf8d5fed | D26 | 0 | 600 | bar |  | 0 | 10 | V | 60 | 0 | [16 empty] | Piezoresistiv | Hydrotechnik GmbH |  | PR 22, 3903-18-33.00 | C6243S | 1 kHz | Rexer | Hydropulser Schrank
0184ebd9-988b-7bb9-b11d-2e53efed30f9 | D27 | 0 | 600 | bar |  | 0 | 10 | V | 60 | 0 | [16 empty] | Piezoresistiv | Hydrotechnik GmbH |  | PR 22, 3903-18-33.00 | S8498S | 1 kHz | Rexer | Hydropulser Schrank
0184ebd9-988b-7bb9-b14d-64bd3544d6b9 | D28 | 0 | 600 | bar |  | 0 | 10 | V | 60 | 0 | [16 empty] | Piezoresistiv | Hydrotechnik GmbH |  | PR 15, 3403-18-71.33 | 1191 | 1 kHz | Rexer | Hydropulser Schrank
0184ebd9-988b-7bb9-b196-99fb4381823b | D29 | nicht erkennbar | nicht erkennbar | bar |  | 0 | 10 | V | #VALUE! | #VALUE! | [16 empty] | DMS | Burster |  | n.e. | n.e. |  | Rexer | Hydropulser Schrank
0184ebd9-988b-7bb9-b1f1-22eed49da9fb | D30 | nicht erkennbar | nicht erkennbar | bar |  | 0 | 10 | V | #VALUE! | #VALUE! | [16 empty] | Piezoresistiv | Keller |  | n.e. | 73152 | 1 kHz | Rexer | Hydropulser Schrank
0184ebd9-988b-7bb9-b200-4f27e1e4ecea | D31 | nicht erkennbar | nicht erkennbar | bar |  | 0 | 10 | V | #VALUE! | #VALUE! | [16 empty] | Piezoresistiv | Keller |  | n.e. | 73153 | 1 kHz | Rexer | Hydropulser Schrank
0184ebd9-988b-7bb9-b26a-1788b6d140c2 | D32 | nicht erkennbar | nicht erkennbar | bar |  | 0 | 10 | V | #VALUE! | #VALUE! | [16 empty] | Piezoresistiv | Keller |  | n.e. | 74312 | 1 kHz | Rexer | Hydropulser Schrank
0184ebd9-988b-7bb9-b297-640d452ef4ba | D33 | nicht erkennbar | nicht erkennbar | bar |  | 0 | 10 | V | #VALUE! | #VALUE! | [16 empty] | Piezoresistiv | Keller |  | n.e. | 88660 | 1 kHz | Rexer | Hydropulser Schrank
0184ebd9-988b-7bb9-bcf2-d23e360009bd | D72 | 0 | 3 | bar | absolut | 0 | 10 | V | 0.3 | 0 | UNKNOWN | UNKNOWN | UNKNOWN | UNKNOWN | 0.00075 | bar | **Reference/relative** pressure sensor: 1 mbar or 0.05 %FS; **Absolute** pressure sensor: 0.5 mbar or 0.025 %FS (10…40 °C) | absolute | 0.003 | bar | <= 0.1% FS | absolute | UNKNOWN | UNKNOWN | UNKNOWN | UNKNOWN | Piezoresistiv | Keller |  | PAA-33X/80794 | 148835 |  | Rexer | Hydropulser Schrank | [5 empty] | S:\Rexer\02_Datenblaetter\datenblatt_150900_Keller_Drucksensor_33x+35x.pdf
0184ebd9-988b-7bb9-bd33-2f48f7a868f2 | D73 | 0 | 10 | bar | absolut | 0 | 10 | V | 1 | 0 | 0.005 | %MV/°C | <= typical: 0.005 %MV/°C (measurement value), maximal: 0.02 %MV/°C (reference temperature according to datasheet most probably at 25°C) | relative_temperature_ambient;relative_mv | 0.0005 | bar/°C | <= typical: 0.005 %FS/°C, maximal: 0.01 %FS/°C (reference temperature according to datasheet most probably at 25°C) | relative_temperature_ambient;absolute | 0.02 | bar | typical: <= +- 0.2 %FS, maximal: <= +- 0.5 %FS | absolute | 0.02 | bar | typical: <= +- 0.2 %FS, maximal: <= +- 0.5 %FS | absolute | Piezoresistiv | Keller |  | PAA-23/8465-10 | 5154 |  | Rexer | Hydropulser Schrank | [5 empty] | S:\Rexer\02_Datenblaetter\datenblatt_080300_Keller_Drucksensor_Serie23u25_200bar.pdf
0184ebd9-988b-7bb9-bd7a-fe02eb1e3573 | D74 | 0 | 100 | bar | absolut | 0 | 10 | V | 10 | 0 | 0.005 | %MV/°C | <= typical: 0.005 %MV/°C (measurement value), maximal: 0.02 %MV/°C (reference temperature according to datasheet most probably at 25°C) | relative_temperature_ambient;relative_mv | 0.005 | bar/°C | <= typical: 0.005 %FS/°C, maximal: 0.01 %FS/°C (reference temperature according to datasheet most probably at 25°C) | relative_temperature_ambient;absolute | 0.2 | bar | typical: <= +- 0.2 %FS, maximal: <= +- 0.5 %FS | absolute | 0.2 | bar | typical: <= +- 0.2 %FS, maximal: <= +- 0.5 %FS | absolute | Piezoresistiv | Keller |  | PA-23/8465-100 | 25194 |  | Rexer | Hydropulser Schrank | [5 empty] | S:\Rexer\02_Datenblaetter\datenblatt_080300_Keller_Drucksensor_Serie23u25_200bar.pdf
0184ebd9-988b-7bb9-bdb4-4125507431a5 | D75 | 0 | 400 | bar | absolut | 0 | 10 | V | 40 | 0 | 0.005 | %MV/°C | <= typical: 0.005 %MV/°C (measurement value), maximal: 0.02 %MV/°C (reference temperature according to datasheet most probably at 25°C) | relative_temperature_ambient;relative_mv | 0.02 | bar/°C | <= typical: 0.005 %FS/°C, maximal: 0.01 %FS/°C (reference temperature according to datasheet most probably at 25°C) | relative_temperature_ambient;absolute | 0.8 | bar | typical: <= +- 0.2 %FS, maximal: <= +- 0.5 %FS | absolute | 0.8 | bar | typical: <= +- 0.2 %FS, maximal: <= +- 0.5 %FS | absolute | Piezoresistiv | Keller |  | PA-23 400 bar | 74909 |  | Rexer | Hydropulser Schrank | [5 empty] | S:\Rexer\02_Datenblaetter\datenblatt_080300_Keller_Drucksensor_Serie23u25_200bar.pdf
0184ebd9-988b-7bb9-bddb-9e34ed0ac9d4 | D76 | 0 | 300 | bar | absolut | 0 | 10 | V | 30 | 0 | 0.005 | %MV/°C | <= typical: 0.005 %MV/°C (measurement value), maximal: 0.02 %MV/°C (reference temperature according to datasheet most probably at 25°C) | relative_temperature_ambient;relative_mv | 0.015 | bar/°C | <= typical: 0.005 %FS/°C, maximal: 0.01 %FS/°C (reference temperature according to datasheet most probably at 25°C) | relative_temperature_ambient;absolute | 0.6 | bar | typical: <= +- 0.2 %FS, maximal: <= +- 0.5 %FS | absolute | 0.6 | bar | typical: <= +- 0.2 %FS, maximal: <= +- 0.5 %FS | absolute | Piezoresistiv | Keller |  | PA-23/8465-300 | 89574 |  | Rexer | Hydropulser Schrank | [5 empty] | S:\Rexer\02_Datenblaetter\datenblatt_080300_Keller_Drucksensor_Serie23u25_200bar.pdf
0184ebd9-988b-7bb9-be00-6e332aae8454 | D077 | 0 | 400 | bar | absolut | 0 | 10 | V | 40 | 0 | 0.005 | %MV/°C | <= typical: 0.005 %MV/°C (measurement value), maximal: 0.02 %MV/°C (reference temperature according to datasheet most probably at 25°C) | relative_temperature_ambient;relative_mv | 0.02 | bar/°C | <= typical: 0.005 %FS/°C, maximal: 0.01 %FS/°C (reference temperature according to datasheet most probably at 25°C) | relative_temperature_ambient;absolute | 0.8 | bar | typical: <= +- 0.2 %FS, maximal: <= +- 0.5 %FS | absolute | 0.8 | bar | typical: <= +- 0.2 %FS, maximal: <= +- 0.5 %FS | absolute | Piezoresistiv | Keller |  | PA-23 400 bar | 74908 |  | Rexer | Hydropulser Schrank | [5 empty] | S:\Rexer\02_Datenblaetter\datenblatt_080300_Keller_Drucksensor_Serie23u25_200bar.pdf
0184ebd9-988b-7bb9-be63-cc1d93cd7036 | D078 | 0 | 400 | bar | absolut | 0 | 10 | V | 40 | 0 | 0.005 | %MV/°C | <= typical: 0.005 %MV/°C (measurement value), maximal: 0.02 %MV/°C (reference temperature according to datasheet most probably at 25°C) | relative_temperature_ambient;relative_mv | 0.02 | bar/°C | <= typical: 0.005 %FS/°C, maximal: 0.01 %FS/°C (reference temperature according to datasheet most probably at 25°C) | relative_temperature_ambient;absolute | 0.8 | bar | typical: <= +- 0.2 %FS, maximal: <= +- 0.5 %FS | absolute | 0.8 | bar | typical: <= +- 0.2 %FS, maximal: <= +- 0.5 %FS | absolute | Piezoresistiv | Keller |  | PA-23 400bar | 74907 |  | Rexer | Hydropulser Schrank | [5 empty] | S:\Rexer\02_Datenblaetter\datenblatt_080300_Keller_Drucksensor_Serie23u25_200bar.pdf
0184ebd9-988b-7bb9-beb8-93b91fedf6fd | D079 | 0 | 5 | bar | absolut | 0 | 10 | V | 0.5 | 0 | 0.005 | %MV/°C | <= typical: 0.005 %MV/°C (measurement value), maximal: 0.02 %MV/°C (reference temperature according to datasheet most probably at 25°C) | relative_temperature_ambient;relative_mv | 0.00025 | bar/°C | <= typical: 0.005 %FS/°C, maximal: 0.01 %FS/°C (reference temperature according to datasheet most probably at 25°C) | relative_temperature_ambient;absolute | 0.01 | bar | typical: <= +- 0.2 %FS, maximal: <= +- 0.5 %FS | absolute | 0.01 | bar | typical: <= +- 0.2 %FS, maximal: <= +- 0.5 %FS | absolute | Piezoresistiv | Keller |  | PAA-23/8465-5 | 88625 |  | Rexer | Hydropulser Schrank | [5 empty] | S:\Rexer\02_Datenblaetter\datenblatt_080300_Keller_Drucksensor_Serie23u25_200bar.pdf
0184ebd9-988b-7bb9-bee8-212e29f8ffb3 | D080 | 0 | 300 | bar | absolut | 0 | 10 | V | 30 | 0 | 0.005 | %MV/°C | <= typical: 0.005 %MV/°C (measurement value), maximal: 0.02 %MV/°C (reference temperature according to datasheet most probably at 25°C) | relative_temperature_ambient;relative_mv | 0.015 | bar/°C | <= typical: 0.005 %FS/°C, maximal: 0.01 %FS/°C (reference temperature according to datasheet most probably at 25°C) | relative_temperature_ambient;absolute | 0.6 | bar | typical: <= +- 0.2 %FS, maximal: <= +- 0.5 %FS | absolute | 0.6 | bar | typical: <= +- 0.2 %FS, maximal: <= +- 0.5 %FS | absolute | Piezoresistiv | Keller |  | PA-23/8465-300 | 89576 |  | Rexer | Hydropulser Schrank | [5 empty] | S:\Rexer\02_Datenblaetter\datenblatt_080300_Keller_Drucksensor_Serie23u25_200bar.pdf
0184ebd9-988b-7bb9-bf0b-8ac992cecf10 | D081 | 0 | 30 | bar | absolut | 0 | 10 | V | 3 | 0 | UNKNOWN | UNKNOWN | UNKNOWN | UNKNOWN | 0.0075 | bar | **Reference/relative** pressure sensor: 1 mbar or 0.05 %FS; **Absolute** pressure sensor: 0.5 mbar or 0.025 %FS (10…40 °C) | absolute | 0.03 | bar | <= 0.1% FS | absolute | UNKNOWN | UNKNOWN | UNKNOWN | UNKNOWN | Piezoresistiv | Keller |  | PAA-33X/80794 | 428736 |  | Rexer | Hydropulser Schrank | [5 empty] | S:\Rexer\02_Datenblaetter\datenblatt_150900_Keller_Drucksensor_33x+35x.pdf
0184ebd9-988b-7bb9-bf54-2e5e70e4e45e | D082 | 0 | 10 | bar | absolut | 0 | 10 | V | 1 | 0 | UNKNOWN | UNKNOWN | UNKNOWN | UNKNOWN | 0.0025 | bar | **Reference/relative** pressure sensor: 1 mbar or 0.05 %FS; **Absolute** pressure sensor: 0.5 mbar or 0.025 %FS (10…40 °C) | absolute | 0.01 | bar | <= 0.1% FS | absolute | UNKNOWN | UNKNOWN | UNKNOWN | UNKNOWN | Piezoresistiv | Keller |  | PAA-33X/80794 | 605587 |  | Rexer | Hydropulser Schrank | [5 empty] | S:\Rexer\02_Datenblaetter\datenblatt_150900_Keller_Drucksensor_33x+35x.pdf
0184ebd9-988b-7bb9-bf88-fa0059d5a748 | D083 | 0 | 30 | bar | absolut | 0 | 10 | V | 3 | 0 | UNKNOWN | UNKNOWN | UNKNOWN | UNKNOWN | 0.0075 | bar | **Reference/relative** pressure sensor: 1 mbar or 0.05 %FS; **Absolute** pressure sensor: 0.5 mbar or 0.025 %FS (10…40 °C) | absolute | 0.03 | bar | <= 0.1% FS | absolute | UNKNOWN | UNKNOWN | UNKNOWN | UNKNOWN | Piezoresistiv | Keller |  | PAA-33X/80794 | 353488 |  | Rexer | Hydropulser Schrank | [5 empty] | S:\Rexer\02_Datenblaetter\datenblatt_150900_Keller_Drucksensor_33x+35x.pdf
0184ebd9-988b-7bb9-bfd8-2298d75171a2 | D084 | 0 | 300 | bar | absolut | 0 | 10 | V | 30 | 0 | UNKNOWN | UNKNOWN | UNKNOWN | UNKNOWN | 0.075 | bar | **Reference/relative** pressure sensor: 1 mbar or 0.05 %FS; **Absolute** pressure sensor: 0.5 mbar or 0.025 %FS (10…40 °C) | absolute | 0.3 | bar | <= 0.1% FS | absolute | UNKNOWN | UNKNOWN | UNKNOWN | UNKNOWN | Piezoresistiv | Keller |  | PAA-33X/300bar | 1090056 |  | Rexer | Hydropulser Schrank | [5 empty] | S:\Rexer\02_Datenblaetter\datenblatt_150900_Keller_Drucksensor_33x+35x.pdf
0184ebd9-988b-7bba-8002-afd53c8a4f26 | D085 | 0 | 300 | bar | absolut | 0 | 10 | V | 30 | 0 | UNKNOWN | UNKNOWN | UNKNOWN | UNKNOWN | 0.075 | bar | **Reference/relative** pressure sensor: 1 mbar or 0.05 %FS; **Absolute** pressure sensor: 0.5 mbar or 0.025 %FS (10…40 °C) | absolute | 0.3 | bar | <= 0.1% FS | absolute | UNKNOWN | UNKNOWN | UNKNOWN | UNKNOWN | Piezoresistiv | Keller |  | PAA-33X/300bar | 1090057 |  | Rexer | Hydropulser Schrank | [5 empty] | S:\Rexer\02_Datenblaetter\datenblatt_150900_Keller_Drucksensor_33x+35x.pdf
0184ebd9-988b-7bba-8051-e4af841965f1 | D086 | 0 | 5 | bar | absolut | 0 | 10 | V | 0.5 | 0 | 0.005 | %MV/°C | <= typical: 0.005 %MV/°C (measurement value), maximal: 0.02 %MV/°C (reference temperature according to datasheet most probably at 25°C) | relative_temperature_ambient;relative_mv | 0.00025 | bar/°C | <= typical: 0.005 %FS/°C, maximal: 0.01 %FS/°C (reference temperature according to datasheet most probably at 25°C) | relative_temperature_ambient;absolute | 0.01 | bar | typical: <= +- 0.2 %FS, maximal: <= +- 0.5 %FS | absolute | 0.01 | bar | typical: <= +- 0.2 %FS, maximal: <= +- 0.5 %FS | absolute | Piezoresistiv | Keller |  | PAA-23/5bar/8465 | 58086 |  | Rexer | Lager Messtechnik | [5 empty] | S:\Rexer\02_Datenblaetter\datenblatt_080300_Keller_Drucksensor_Serie23u25_200bar.pdf
0184ebd9-988b-7bba-8089-4ba1eecb242e | D087 | 0 | 4 | bar | absolut | 0 | 10 | V | 0.4 | 0 | 0.005 | %MV/°C | <= typical: 0.005 %MV/°C (measurement value), maximal: 0.02 %MV/°C (reference temperature according to datasheet most probably at 25°C) | relative_temperature_ambient;relative_mv | 0.0002 | bar/°C | <= typical: 0.005 %FS/°C, maximal: 0.01 %FS/°C (reference temperature according to datasheet most probably at 25°C) | relative_temperature_ambient;absolute | 0.008 | bar | typical: <= +- 0.2 %FS, maximal: <= +- 0.5 %FS | absolute | 0.008 | bar | typical: <= +- 0.2 %FS, maximal: <= +- 0.5 %FS | absolute | Piezoresistiv | Keller |  | PAA-23/8465.1-4 | 22588 |  | Rexer | Lager Messtechnik | [5 empty] | S:\Rexer\02_Datenblaetter\datenblatt_080300_Keller_Drucksensor_Serie23u25_200bar.pdf
0184ebd9-988b-7bba-813a-824dda198d3b | D088 | -10 | 10 | bar | absolut | 0 | 10 | V | 2 | -10 | [16 empty] | Differenzdruck | Keller |  | PD-23/8666- -10/+10 | 79699 |  | Rexer | Lager Messtechnik | [5 empty] | S:\Rexer\02_Datenblaetter\datenblatt_080300_Keller_Drucksensor_Serie23u25_200bar.pdf
0184ebd9-988b-7bba-815a-cd1b41f3fba6 | D089 | -20 | 20 | bar | absolut | 0 | 10 | V | 4 | -20 | [16 empty] | Differenzdruck | Keller |  | PD-23/8666 -20/+20 | 74999 |  | Rexer | Lager Messtechnik | [5 empty] | S:\Rexer\02_Datenblaetter\datenblatt_080300_Keller_Drucksensor_Serie23u25_200bar.pdf
0184ebd9-988b-7bba-81be-82f24102228f | D090 | 0 | 200 | bar | relativ | 0 | 10 | V | 20 | 0 | [16 empty] | Piezoresistiv | Sensortechnics |  | KTE56200GQ4 |  |  | Rexer | Lager Messtechnik
0184ebd9-988b-7bba-81ff-5c01dcec45ba | D091 | 0 | 200 | bar | relativ | 0 | 10 | V | 20 | 0 | [16 empty] | Piezoresistiv | Sensortechnics |  | KTE56200GQ4 |  |  | Rexer | Lager Messtechnik
0184ebd9-988b-7bba-8203-06be5cf6bbb8 | D092 | 0 | 10 | bar | absolut | 0 | 10 | V | 1 | 0 | UNKNOWN | UNKNOWN | UNKNOWN | UNKNOWN | 0.0025 | bar | **Reference/relative** pressure sensor: 1 mbar or 0.05 %FS; **Absolute** pressure sensor: 0.5 mbar or 0.025 %FS (10…40 °C) | absolute | 0.01 | bar | <= 0.1% FS | absolute | UNKNOWN | UNKNOWN | UNKNOWN | UNKNOWN | Piezoresistiv | Keller |  | PAA-33X/10bar | 1011246 |  | Rexer | Hydropulser Schrank | [5 empty] | S:\Rexer\02_Datenblaetter\datenblatt_150900_Keller_Drucksensor_33x+35x.pdf
0184ebd9-988b-7bba-8271-d7a0eacb5efd | D093 | 0 | 10 | bar | absolut | 0 | 10 | V | 1 | 0 | UNKNOWN | UNKNOWN | UNKNOWN | UNKNOWN | 0.0025 | bar | **Reference/relative** pressure sensor: 1 mbar or 0.05 %FS; **Absolute** pressure sensor: 0.5 mbar or 0.025 %FS (10…40 °C) | absolute | 0.01 | bar | <= 0.1% FS | absolute | UNKNOWN | UNKNOWN | UNKNOWN | UNKNOWN | Piezoresistiv | Keller |  | PAA-33X/10bar | 1011233 |  | Rexer | Hydropulser Schrank | [5 empty] | S:\Rexer\02_Datenblaetter\datenblatt_150900_Keller_Drucksensor_33x+35x.pdf
0184ebd9-988b-7bba-828c-2a01abd07824 | D094 | 1 | 201 | bar | relativ | 0 | 10 | V | 20 | 1 | 0.005 | %MV/°C | <= typical: 0.005 %MV/°C (measurement value), maximal: 0.02 %MV/°C (reference temperature according to datasheet most probably at 25°C) | relative_temperature_ambient;relative_mv | 0.01 | bar/°C | <= typical: 0.005 %FS/°C, maximal: 0.01 %FS/°C (reference temperature according to datasheet most probably at 25°C) | relative_temperature_ambient;absolute | 0.4 | bar | typical: <= +- 0.2 %FS, maximal: <= +- 0.5 %FS | absolute | 0.4 | bar | typical: <= +- 0.2 %FS, maximal: <= +- 0.5 %FS | absolute | Piezoresistiv | Keller |  | PA23/8465-200 | 88662 |  | Rexer | Hydropulser Prüfstand | [5 empty] | S:\Rexer\02_Datenblaetter\datenblatt_080300_Keller_Drucksensor_Serie23u25_200bar.pdf
0184ebd9-988b-7bba-82f2-a3655161db0d | D095 | 1 | 201 | bar | relativ | 0 | 10 | V | 20 | 1 | 0.005 | %MV/°C | <= typical: 0.005 %MV/°C (measurement value), maximal: 0.02 %MV/°C (reference temperature according to datasheet most probably at 25°C) | relative_temperature_ambient;relative_mv | 0.01 | bar/°C | <= typical: 0.005 %FS/°C, maximal: 0.01 %FS/°C (reference temperature according to datasheet most probably at 25°C) | relative_temperature_ambient;absolute | 0.4 | bar | typical: <= +- 0.2 %FS, maximal: <= +- 0.5 %FS | absolute | 0.4 | bar | typical: <= +- 0.2 %FS, maximal: <= +- 0.5 %FS | absolute | Piezoresistiv | Keller |  | PA23/8465-200 | 25186 |  | Rexer | Hydropulser Prüfstand | [5 empty] | S:\Rexer\02_Datenblaetter\datenblatt_080300_Keller_Drucksensor_Serie23u25_200bar.pdf
0184ebd9-988b-7bba-8310-d3fc5137ddf6 | D096 | 1 | 201 | bar | relativ | 0 | 10 | V | 20 | 1 | 0.005 | %MV/°C | <= typical: 0.005 %MV/°C (measurement value), maximal: 0.02 %MV/°C (reference temperature according to datasheet most probably at 25°C) | relative_temperature_ambient;relative_mv | 0.01 | bar/°C | <= typical: 0.005 %FS/°C, maximal: 0.01 %FS/°C (reference temperature according to datasheet most probably at 25°C) | relative_temperature_ambient;absolute | 0.4 | bar | typical: <= +- 0.2 %FS, maximal: <= +- 0.5 %FS | absolute | 0.4 | bar | typical: <= +- 0.2 %FS, maximal: <= +- 0.5 %FS | absolute | Piezoresistiv | Keller |  | PA23/8465-200 | 25185 |  | Rexer | Hydropulser Prüfstand | [5 empty] | S:\Rexer\02_Datenblaetter\datenblatt_080300_Keller_Drucksensor_Serie23u25_200bar.pdf
0184ebd9-988b-7bba-8365-2c49aaa69770 | D097 | 0 | 200 | bar | absolut | 0 | 10 | V | 20 | 0 | 0.005 | %MV/°C | <= typical: 0.005 %MV/°C (measurement value), maximal: 0.02 %MV/°C (reference temperature according to datasheet most probably at 25°C) | relative_temperature_ambient;relative_mv | 0.01 | bar/°C | <= typical: 0.005 %FS/°C, maximal: 0.01 %FS/°C (reference temperature according to datasheet most probably at 25°C) | relative_temperature_ambient;absolute | 0.4 | bar | typical: <= +- 0.2 %FS, maximal: <= +- 0.5 %FS | absolute | 0.4 | bar | typical: <= +- 0.2 %FS, maximal: <= +- 0.5 %FS | absolute | Piezoresistiv | Keller |  | PA23 200bar |  |  | Rexer | Hydropulser Prüfstand | [5 empty] | S:\Rexer\02_Datenblaetter\datenblatt_080300_Keller_Drucksensor_Serie23u25_200bar.pdf
0184ebd9-988b-7bba-83a5-01cec15c9820 | D098 | 0 | 50 | bar | absolut | 0 | 10 | V | 5 | 0 | UNKNOWN | UNKNOWN | UNKNOWN | UNKNOWN | 0.0125 | bar | **Reference/relative** pressure sensor: 1 mbar or 0.05 %FS; **Absolute** pressure sensor: 0.5 mbar or 0.025 %FS (10…40 °C) | absolute | 0.05 | bar | <= 0.1% FS | absolute | UNKNOWN | UNKNOWN | UNKNOWN | UNKNOWN | Piezoresistiv | Keller |  | PAA-33X/80794 | 431003 |  | Rexer | Hydropulser Prüfstand | [5 empty] | S:\Rexer\02_Datenblaetter\datenblatt_150900_Keller_Drucksensor_33x+35x.pdf
0184ebd9-988b-7bba-83e4-e24d90bf1352 | D099 | 0 | 50 | bar | absolut | 0 | 10 | V | 5 | 0 | UNKNOWN | UNKNOWN | UNKNOWN | UNKNOWN | 0.0125 | bar | **Reference/relative** pressure sensor: 1 mbar or 0.05 %FS; **Absolute** pressure sensor: 0.5 mbar or 0.025 %FS (10…40 °C) | absolute | 0.05 | bar | <= 0.1% FS | absolute | UNKNOWN | UNKNOWN | UNKNOWN | UNKNOWN | Piezoresistiv | Keller |  | PAA-33X/80794 | 535xxx |  | Rexer | Hydropulser Prüfstand | [5 empty] | S:\Rexer\02_Datenblaetter\datenblatt_150900_Keller_Drucksensor_33x+35x.pdf
0184ebd9-988b-7bba-8402-16963d276842 | D100 | 0 | 30 | bar | absolut | 0 | 10 | V | 3 | 0 | [16 empty] | Piezoresistiv | Keller |  | PAA-35X/80797 | 126240 |  | Rexer | Lager Messtechnik
018a0584-99c5-7e34-9386-8d7897b664e4 | D129 | 0 | 10 | bar | absolut | 0 | 10 | V | 1 | 1 | UNKNOWN | UNKNOWN | UNKNOWN | UNKNOWN | 0.04 | bar | <0.5% FS | absolute | 0.025 | bar | <0.25% FS | absolute | UNKNOWN | UNKNOWN | UNKNOWN | UNKNOWN |  | WIKA |  | A-10 | 1A028EUETJ1 |  | Rexer | [5 empty] | neu
018a058b-c08f-7c04-a5ac-eb91780ae4aa | D130 | 0 | 10 | bar | relativ | 0 | 10 | V | 1 | 1 | UNKNOWN | UNKNOWN | UNKNOWN | UNKNOWN | 0.05 | bar | <0.5% FS | absolute | 0.05 | bar | <0.5% FS | absolute | UNKNOWN | UNKNOWN | UNKNOWN | UNKNOWN |  | WIKA |  | O-10 (C) | 110J2CXA |  | Rexer | [5 empty] | neu
018a058b-c090-7ea8-af82-344c238f1077 | D131 | 0 | 10 | bar | absolut | 0 | 10 | V | 1 | 1 | UNKNOWN | UNKNOWN | UNKNOWN | UNKNOWN | 0.04 | bar | <0.5% FS | absolute | 0.025 | bar | <0.25% FS | absolute | UNKNOWN | UNKNOWN | UNKNOWN | UNKNOWN |  | WIKA |  | A-10 | 1A028EU6625 |  | Rexer | [5 empty] | neu
018a058b-c091-7d2e-a809-52a2e0fb9dc6 | D132 | 0 | 10 | bar | relativ | 0 | 10 | V | 1 | 1 | UNKNOWN | UNKNOWN | UNKNOWN | UNKNOWN | 0.05 | bar | <0.5% FS | absolute | 0.05 | bar | <0.5% FS | absolute | UNKNOWN | UNKNOWN | UNKNOWN | UNKNOWN |  | WIKA |  | O-10 (T) | 110J2CYM |  | Rexer | [5 empty] | neu

## Kraft
uuid | Ident-Nummer | Messbereich von | Messbereich bis | Messbereich Einheit | Ausgabebereich von | Ausgabebereich bis | Ausgabebereich Einheit | Kennlinie Steigung _ Sensitivity | Kennlinie Offset _ Bias | Sensitivity Uncertainty | Sensitivity Uncertainty Unit | Sensitivity Uncertainty Comment | Sensitivity Uncertainty Keywords | Bias Uncertainty | Bias Uncertainty Unit | Bias Uncertainty Comment | Bias Uncertainty Keywords | Linearity Uncertainty | Linearity Uncertainty Unit | Linearity Uncertainty Comment | Linearity Uncertainty Keywords | Hysteresis Uncertainty | Hysteresis Uncertainty Unit | Hysteresis Uncertainty Comment | Hysteresis Uncertainty Keywords | Messprinzip | Hersteller | Hersteller P_ID | Bezeichnung | Seriennummer | Verantwortlicher WiMi | Aufbewahrungsort | Datenblatt Link | Kalibrationsprotokoll Link | letzte Prüfung/ Kalibration | Zubehör | Bemerkung | Datasheet FST-Link
0184ebd9-988c-7bbb-9599-8d7f6ace9cb8 | K01 | 0 | 5 | kN | 0 | 10 | V | 0.5 | 0 | [16 empty] | DMS | MTS |  | 661.19F-01 | 369955 | Rexer | Hydropulser Schrank
0184ebd9-988c-7bbb-95cd-c16f369c5d01 | K02 | 0 | 10 | kN | 0 | 10 | V | 1 | 0 | [16 empty] | DMS | HBM |  | 1-S9M/10KN-1 | 30893194 | Rexer | Hydropulser Schrank
0184ebd9-988c-7bbb-9617-8e614e5f0e4f | K03 | 0 | 50 | kN | 0 | 10 | V | 5 | 0 | [16 empty] | DMS | MTS |  | 661.20F-02 | 356069 | Rexer | Hydropulser Schrank
0184ebd9-988c-7bbb-964d-711052667695 | K04 | 0 | 10 | kN | 0 | 10 | V | 1 | 0 | [16 empty] | DMS | HBM |  | 1-C2/10kN | 1000/148893 | Rexer | Hydropulser Schrank
0184ebd9-988c-7bbb-96a8-78c88744acf2 | K05 | 0 | 10 | kN | 0 | 10 | V | 1 | 0 | [16 empty] | DMS | HBM |  | 1-C2/10kN | 1000/155416 | Rexer | Hydropulser Schrank
0184ebd9-988c-7bbb-96c9-73a6972fe09e | K06 | 0 | 10 | kN | 0 | 10 | V | 1 | 0 | [16 empty] | DMS | HBM |  | U2A | D29507 1t | Rexer | Hydropulser Schrank
0184ebd9-988c-7bbb-9770-edb6ba36b617 | K07 | 0 | 50 | kN | 0 | 10 | V | 5 | 0 | 0.015 | %/10K | <=0.015% Temperatureinfluss auf den kennwert/10K auf dne Kennwert ausserhalb des Nentemperatursbereich (-10°C bis 45°C) | relative_temperature_ambient; relative_mv | 0.5 | kN | <= 1% FS | absolute | 0.025 | kN | <= 0.05 %FS | absolute | 0.025 | kN | < 0.125% at 0.4 FS | absolute | DMS | HBM |  | U10M/50kN | 014879S | Rexer | Hydropulser Schrank | [5 empty] | S:\Rexer\02_Datenblaetter\datenblatt_000000_HBM_Kraftsensor_U10M.pdf
0184ebd9-988c-7bbb-9794-3dce416ae94d | K08 | 0 | 25 | kN | 0 | 10 | V | 2.5 | 0 | 0.015 | %/10K | <=0.015% Temperatureinfluss auf den kennwert/10K auf dne Kennwert ausserhalb des Nentemperatursbereich (-10°C bis 45°C) | relative_temperature_ambient; relative_mv | 0.25 | kN | <= 1% FS | absolute | 0.01 | kN | <= 0.04% FS | absolute | 0.01 | kN | < 0.1% at 0.4 FS | absolute | DMS | HBM |  | U10M/25kN | 019581S | Rexer | Hydropulser Schrank | [5 empty] | S:\Rexer\02_Datenblaetter\datenblatt_000000_HBM_Kraftsensor_U10M.pdf
0184ebd9-988c-7bbb-97cd-63e191335f56 | K09 | 0 | 5 | kN | 0 | 10 | V | 0.5 | 0 | 0.015 | %/10K | <=0.015% Temperatureinfluss auf den kennwert/10K auf dne Kennwert ausserhalb des Nentemperatursbereich (-10°C bis 45°C) | relative_temperature_ambient; relative_mv | 0.05 | kN | <= 1% FS | absolute | 0.0015 | kN | <= 0.03% FS | absolute | 0.0015 | kN | < 0.075% at 0.4 FS | absolute | DMS | HBM |  | U10M/5kN | 022726S | Rexer | Hydropulser Schrank | [5 empty] | S:\Rexer\02_Datenblaetter\datenblatt_000000_HBM_Kraftsensor_U10M.pdf
0184ebd9-988c-7bbb-9802-abdb21f0289c | K10 | 0 | 50 | kN | 0 | 10 | V | 5 | 0 | [16 empty] | DMS | Althen |  | ALF300UFR0K0 | 50432 | Rexer | Hydropulser Schrank
0184ebd9-988c-7bbb-9870-4fa72eb37059 | K11 | 0 | 12.5 | kN | 0 | 10 | V | 1.25 | 0 | 0.015 | %/10K | <=0.015% Temperatureinfluss auf den kennwert/10K auf dne Kennwert ausserhalb des Nentemperatursbereich (-10°C bis 45°C) | relative_temperature_ambient; relative_mv | 0.125 | kN | <= 1% FS | absolute | 0.005 | kN | <= 0.04% FS | absolute | 0.005 | kN | < 0.1% at 0.4 FS | absolute | DMS | HBM |  | U10M/12,5kN | 016539S | Rexer | Hydropulser Prüfstand | [5 empty] | S:\Rexer\02_Datenblaetter\datenblatt_000000_HBM_Kraftsensor_U10M.pdf

## Weg
uuid | Ident-Nummer | Messbereich von | Messbereich bis | Messbereich Einheit | Ausgabebereich von | Ausgabebereich bis | Ausgabebereich Einheit | Kennlinie Steigung _ Sensitivity | Kennlinie Offset _ Bias | Sensitivity Uncertainty | Sensitivity Uncertainty Unit | Sensitivity Uncertainty Comment | Sensitivity Uncertainty Keywords | Bias Uncertainty | Bias Uncertainty Unit | Bias Uncertainty Comment | Bias Uncertainty Keywords | Linearity Uncertainty | Linearity Uncertainty Unit | Linearity Uncertainty Comment | Linearity Uncertainty Keywords | Hysteresis Uncertainty | Hysteresis Uncertainty Unit | Hysteresis Uncertainty Comment | Hysteresis Uncertainty Keywords | Messprinzip | Hersteller | Hersteller P_ID | Bezeichnung | Seriennummer | Verantwortlicher WiMi | Aufbewahrungsort | Datenblatt Link | Kalibrationsprotokoll Link | letzte Prüfung/ Kalibration | Zubehör | Bemerkung | Datasheet FST-Link
0184ebd9-988c-7bbb-90d4-1ba867e7a4fb | W01 | 0 | 10 | mm | 0 | 10 | V | 1 | 0 | [16 empty] | Induktiv | WayCon |  | SM10-A-SA | 826026 | Rexer | Hydropulser Schrank
0184ebd9-988c-7bbb-9273-0f24db974e80 | W06 | 0 | 25 | mm | 0 | 10 | V | 2.5 | 0 | [16 empty] | Mechanisch | Novotechnik |  | TS-0025 | 130583 | Rexer | Hydropulser Schrank
0184ebd9-988c-7bbb-92b0-b384b573414b | W07 | 0 | 25 | mm | 0 | 10 | V | 2.5 | 0 | [16 empty] | Mechanisch | Novotechnik |  | TS-0025 | 130584 | Rexer | Hydropulser Schrank
0184ebd9-988c-7bbb-92ce-75543625f949 | W08 | 0 | 25 | mm | 0 | 10 | V | 2.5 | 0 | [16 empty] | Mechanisch | Novotechnik |  | TS-0025 | 130581 | Rexer | Hydropulser Schrank
0184ebd9-988c-7bbb-9333-9254e5ab4368 | W09 | 65 | 250 | mm | 0 | 10 | V | 18.5 | 65 | [16 empty] | Optisch | RIFTEK |  | RF605 | 25682 | Rexer
0184ebd9-988c-7bbb-9368-c512c9bb6eac | W10 | 65 | 250 | mm | 0 | 10 | V | 18.5 | 65 | [16 empty] | Optisch | RIFTEK |  | RF605 | 25662 | Rexer
0184ebd9-988c-7bbb-93a3-8a97bf0e7983 | W11 | 0 | 150 | mm | 0 | 10 | V | 15 | 0 | [16 empty] | Induktiv | Novotechnik |  | LWG150 |  | Rexer
0184ebd9-988c-7bbb-93da-aa6d26b5e1fc | W12 | 0 | 225 | mm | 0 | 10 | V | 22.5 | 0 | [16 empty] | Induktiv | Novotechnik |  | LWG225 |  | Rexer
0185217b-73cb-750b-abbf-a76c31616b15 | W19 | 0 | 300 | mm | -10 | 10 | V | 15 | 0 | UNKNOWN | UNKNOWN | UNKNOWN | UNKNOWN | UNKNOWN | UNKNOWN | UNKNOWN | UNKNOWN | 1 | % | 1%MV (of the measurement value) according to the datasheet | relative_mv | UNKNOWN | UNKNOWN | UNKNOWN | UNKNOWN | induktiv | MTS |  | LVDT | 039-075-105 | Rexer | Hydropulser |  |  |  |  | Im Hydropulser fest verbaut | S:\Rexer\02_Datenblaetter\datenblatt_000000_MTS_Wegsensor_LVDT.pdf

## Temperatur
uuid | Ident-Nummer | Messbereich von | Messbereich bis | Messbereich Einheit | Ausgabebereich von | Ausgabebereich bis | Ausgabebereich Einheit | Kennlinie Steigung _ Sensitivity | Kennlinie Offset _ Bias | Sensitivity Uncertainty | Sensitivity Uncertainty Unit | Sensitivity Uncertainty Comment | Sensitivity Uncertainty Keywords | Bias Uncertainty | Bias Uncertainty Unit | Bias Uncertainty Comment | Bias Uncertainty Keywords | Linearity Uncertainty | Linearity Uncertainty Unit | Linearity Uncertainty Comment | Linearity Uncertainty Keywords | Hysteresis Uncertainty | Hysteresis Uncertainty Unit | Hysteresis Uncertainty Comment | Hysteresis Uncertainty Keywords | Messprinzip | Hersteller | Hersteller P_ID | Bezeichnung | Seriennummer | Verantwortlicher WiMi | Aufbewahrungsort | Datenblatt Link | Kalibrationsprotokoll Link | letzte Prüfung/ Kalibration | Zubehör | Bemerkung | Datasheet FST-Link
0184ebd9-988c-7bbb-9f6d-fca5fbdfdd0a | T01 | -50 | 600 | °C | 0 | 10 | V | 65 | -50 | [16 empty] | Pt100 | Jumo |  | 90.295-F74, 232137000 | 97450008 | Rexer | Hydropulser Schrank
0184ebd9-988c-7bbb-a2a8-661580f4461c | T13 | -25 | 100 | °C | 0 | 10 | V | 12.5 | -25 | UNKNOWN | UNKNOWN | UNKNOWN | UNKNOWN | 0.0125 | °C/°C | <= 0.01%FS/°C (reference point not closer described, maybe 25°C) | relative_temperature_ambient; absolute | 0.5 | °C | typical: <= +-0.4% FS, maximal: <= +-0.8% FS | absolute | UNKNOWN | UNKNOWN | UNKNOWN | UNKNOWN | Pt1000 | Hydac |  | ETS 4146-B-006-000 | 617P014983 | Rexer | Hydropulser Schrank | [5 empty] | S:\Rexer\02_Datenblaetter\datenblatt_130116_Hydac_Temperatur_Messumformer_ETS_41XX.pdf
0184ebd9-988c-7bbb-a2c0-8ba93f63085c | T14 | -25 | 150 | °C | 0 | 10 | V | 17.5 | -25 | UNKNOWN | UNKNOWN | UNKNOWN | UNKNOWN | 0.5 | °C | <= +-0.5°C (to the reference point) | absolute | 1.75 | °C | < 1% FS (depends on the single thermocouple) | absolute | UNKNOWN | UNKNOWN | UNKNOWN | UNKNOWN | Messumformer/Thermoelement Typ K | LKM |  | K -25..150°C | 163616 | Rexer | Hydropulser Schrank | [5 empty] | S:\Rexer\02_Datenblaetter\datenblatt_110200_LKM_Temperatur_Messumformer Thermoelement_LKM_TYP102.pdf
0184ebd9-988c-7bbb-a319-f28d8782ed65 | T15 | 0 | 100 | °C | 0 | 10 | V | 10 | 0 | UNKNOWN | UNKNOWN | UNKNOWN | UNKNOWN | 0.5 | °C | <= +-0.5°C (to the reference point) | absolute | 1 | °C | < 1% FS (depends on the single thermocouple) | absolute | UNKNOWN | UNKNOWN | UNKNOWN | UNKNOWN | Messumformer/Thermoelement Typ K | LKM |  | K 0..100°C | 243008 | Rexer | Hydropulser Schrank | [5 empty] | S:\Rexer\02_Datenblaetter\datenblatt_110200_LKM_Temperatur_Messumformer Thermoelement_LKM_TYP102.pdf
0184ebd9-988c-7bbb-a350-11bd6316decd | T16 | 0 | 100 | °C | 0 | 10 | V | 10 | 0 | UNKNOWN | UNKNOWN | UNKNOWN | UNKNOWN | 0.5 | °C | <= +-0.5°C (to the reference point) | absolute | 1 | °C | < 1% FS (depends on the single thermocouple) | absolute | UNKNOWN | UNKNOWN | UNKNOWN | UNKNOWN | Messumformer/Thermoelement Typ K | LKM |  | K 0..100°C | 203805 | Rexer | Hydropulser Schrank | [5 empty] | S:\Rexer\02_Datenblaetter\datenblatt_110200_LKM_Temperatur_Messumformer Thermoelement_LKM_TYP102.pdf
0184ebd9-988c-7bbb-a38e-1797d4efe338 | T17 | -25 | 100 | °C | 0 | 10 | V | 12.5 | -25 | UNKNOWN | UNKNOWN | UNKNOWN | UNKNOWN | 0.0125 | °C/°C | <= 0.01%FS/°C (reference point not closer described, maybe 25°C) | relative_temperature_ambient; absolute | 0.5 | °C | typical: <= +-0.4% FS, maximal: <= +-0.8% FS | absolute | UNKNOWN | UNKNOWN | UNKNOWN | UNKNOWN | Pt1000 | Hydac |  | ETS 4146-B-006-000 | 827P022412 | Rexer | Hydropulser Prüfstand | [5 empty] | S:\Rexer\02_Datenblaetter\datenblatt_130116_Hydac_Temperatur_Messumformer_ETS_41XX.pdf
018a058b-c08d-79ba-8e37-2a4c2bec56b5 | T22 | -40 | 100 | °C | 0 | 10 | V | 14 | -40 | UNKNOWN | UNKNOWN | UNKNOWN | UNKNOWN | 0.5 | °C | <= +-0.5°C (to the reference point) | absolute | 1.4 | °C | < 1% FS (depends on the single thermocouple) | absolute | UNKNOWN | UNKNOWN | UNKNOWN | UNKNOWN | Messumformer/Thermoelement Typ K | LKM |  | TYP 102 | 184416 Kennnummer | Rexer | Hydropulser Schrank | [5 empty] | S:\Rexer\02_Datenblaetter\datenblatt_110200_LKM_Temperatur_Messumformer Thermoelement_LKM_TYP102.pdf
018a058b-c08e-740e-b9cc-08953e4952ab | T23 | -40 | 100 | °C | 0 | 10 | V | 14 | -40 | UNKNOWN | UNKNOWN | UNKNOWN | UNKNOWN | 0.5 | °C | <= +-0.5°C (to the reference point) | absolute | 1.4 | °C | < 1% FS (depends on the single thermocouple) | absolute | UNKNOWN | UNKNOWN | UNKNOWN | UNKNOWN | Messumformer/Thermoelement Typ K | LKM |  | TYP 102 | 184416 Kennnummer | Rexer | Hydropulser Schrank | [5 empty] | S:\Rexer\02_Datenblaetter\datenblatt_110200_LKM_Temperatur_Messumformer Thermoelement_LKM_TYP102.pdf
064f05d1-5d2d-7a6f-8000-a3da10f5a1a3 | T24 | -25 | 100 | °C | 0 | 10 | V | 12.5 | -25 | UNKNOWN | UNKNOWN | UNKNOWN | UNKNOWN | 0.0125 | °C/°C | <= 0.01%FS/°C (reference point not closer described, maybe 25°C) | relative_temperature_ambient; absolute | 0.5 | °C | typical: <= +-0.4% FS, maximal: <= +-0.8% FS | absolute | UNKNOWN | UNKNOWN | UNKNOWN | UNKNOWN | Pt1000 | Hydac |  | ETS 4146-B-006-000 | 648P016436 | Rexer | Hydropulser Prüfstand | [5 empty] | S:\Rexer\02_Datenblaetter\datenblatt_130116_Hydac_Temperatur_Messumformer_ETS_41XX.pdf

## Leistung
uuid | Ident-Nummer | Sensor Messbereich von | Sensor Messbereich bis | Sensor Messbereich Einheit | Sensor Ausgabebereich von | Sensor Ausgabebereich bis | Sensor Ausgabebereich Einheit | Sensor Kennlinie _ Sensitivity | Sensor Kennlinie _ Bias | Sensitivity Uncertainty | Sensitivity Uncertainty Unit | Sensitivity Uncertainty Comment | Sensitivity Uncertainty Keywords | Absolute Bias Uncertainty | Absolute Bias Uncertainty Unit | Absolute Bias Uncertainty Comment | Absolute Bias Uncertainty Keywords | Relative Bias Uncertainty | Relative Bias Uncertainty Unit | Relative Bias Uncertainty Comment | Relative Bias Uncertainty Keywords | Linearity Uncertainty | Linearity Uncertainty Unit | Linearity Uncertainty Comment | Linearity Uncertainty Keywords | Hysteresis Uncertainty | Hysteresis Uncertainty Unit | Hysteresis Uncertainty Comment | Hysteresis Uncertainty Keywords | Messprinzip | Hersteller | Hersteller P_ID | Bezeichnung | Seriennummer | Verantwortlicher WiMi | Aufbewahrungsort | Datenblatt Link | Kalibrationsprotokoll Link | letzte Prüfung/ Kalibration | Zubehör | Bemerkung | Datasheet FST-Link

## Volumenstrom
uuid | Ident-Nummer | Messbereich von | Messbereich bis | Messbereich Einheit | Ausgabebereich von | Ausgabebereich bis | Ausgabebereich Einheit | Kennlinie Steigung _ Sensitivity | Kennlinie Offset _ Bias | Sensitivity Uncertainty | Sensitivity Uncertainty Unit | Sensitivity Uncertainty Comment | Sensitivity Uncertainty Keywords | Bias Uncertainty | Bias Uncertainty Unit | Bias Uncertainty Comment | Bias Uncertainty Keywords | Linearity Uncertainty | Linearity Uncertainty Unit | Linearity Uncertainty Comment | Linearity Uncertainty Keywords | Hysteresis Uncertainty | Hysteresis Uncertainty Unit | Hysteresis Uncertainty Comment | Hysteresis Uncertainty Keywords | Messprinzip | Hersteller | Hersteller P_ID | Bezeichnung | Seriennummer | Verantwortlicher WiMi | Aufbewahrungsort | Datenblatt Link | Kalibrationsprotokoll Link | letzte Prüfung/ Kalibration | Zubehör | Bemerkung | Datasheet FST-Link